## Slide 1: Keeping Labor Safe – Phase 2
Machine Learning Outcomes
Lucas Finco

## Slide 2: Analysis Process (to date)
Extract data samples of traces to re-produce strip
Identify UA baseline
Identify contractions
Calculate U, R, and U/R ratio
Calculate FHR baseline
Calculate FHR variability
Identify late recoveries
Identify accelerations

## Slide 3: Fetal Risk Index Components
Excessive UA
FHR Baseline Range
FHR Variability
FHR Accelerations
FHR Late Recovery
Fetal Risk Factor
Maternal Risk Factor
Obstetric Risk Factor

## Slide 4: UA Thresholding
Samples of UA traces extracted for testing
Baseline (Orange)
Threshold (Green)
If mode is too high (above median), we use 25% quantile instead

## Slide 5: UA Contraction Identification
Count up time above Threshold
Clearly indicates Contractions
Mark all < 30sec

## Slide 6: UA Contraction Identification
Contractions (Green)
Note contractions currently start 30sec in (we correct for this in the U/R ratio calculations)
In this example, 3 contractions are identified

## Slide 7: FHR Baseline
Use time not in contraction
Also remove all FHR=0 data
Identify FHR baseline via mode method
Again, if mode is above median, use the 25% quantile instead
Variability is standard deviation over same timespan

## Slide 8: FHR Accelerations
Identify FHR levels more than 10 BPM above FHR baseline for more than 10 seconds
In this example, 1 acceleration is identified (purple line)

## Slide 9: FHR Late Recoveries
Start at the end of each contraction
Identify FHR levels deviating from baseline more than 15 BPM above after 15 seconds
In this example, 5 late recoveries are identified

## Slide 10: Results
Ground Truth provided by Dr. Eden’s review of strips provided
Number of observations: 22

## Slide 11: Results: Excessive UA
Number of contractions exceeds 4 in 10 minutes
Accuracy: 86%
Sensitivity, Recall: 78%
Specificity: 92%
Precision, PPV: 88%
NPV: 86%
True Negative Rate: 92%
Good result, confident in conservative prediction, but could improve make less
 | Actual |  | 
Predicted |  | 0 | 1
 | 0 | 12 | 2
 | 1 | 1 | 7

## Slide 12: Results: FHR Baseline
FHR Baseline < 110 or >= 160
Accuracy: 100%
Sensitivity, Recall: 100%
Specificity: 100%
Precision, PPV: 100%
NPV: 100%
True Negative Rate: 100%
Very strong result, note only one zero, no improvement needed
 | Actual |  | 
Predicted |  | 0 | 1
 | 0 | 1 | 0
 | 1 | 0 | 21

## Slide 13: Results: FHR Accelerations
FHR Accelerations = 0
Accuracy: 91%
Sensitivity, Recall: 100%
Specificity: 67%
Precision, PPV: 89%
NPV: 100%
True Negative Rate: 67%
Good result, zero false negatives, but 2 false positives
 | Actual |  | 
Predicted |  | 0 | 1
 | 0 | 4 | 0
 | 1 | 2 | 16

## Slide 14: Results: FHR Variability
FHR Variability < 5 or > 15
Accuracy: 77%
Sensitivity, Recall: 92%
Specificity: 60%
Precision, PPV: 73%
NPV: 86%
True Negative Rate: 60%
Good result, better at predicting zeros, which is important, but missed 5 zeros
 | Actual |  | 
Predicted |  | 0 | 1
 | 0 | 6 | 1
 | 1 | 4 | 11

## Slide 15: Results: FHR Late Recovery
FHR Late Recovery Present
Accuracy: 77%
Sensitivity, Recall: 88%
Specificity: 50%
Precision, PPV: 82%
NPV: 60%
True Negative Rate: 50%
Good result, confident in prediction, but weak with actual zeros, and 2 false negatives
 | Actual |  | 
Predicted |  | 0 | 1
 | 0 | 3 | 2
 | 1 | 3 | 14

## Slide 16: ToDo’s:
Use AUC method for late recoveries
Overshoots
Decelerations
Data Quality Metric
Analysis Confidence Metric

## Slide 17: Recommendations
Collect ground truth at the source of the data
  When contractions happen
  When accelerations happen
  When late recoveries happen

## Slide 18: Appendix

## Slide 19: UA Curve Fitting
Samples of UA traces extracted and fit for testing
Fitting normal curve y = a +b * ec*t2

## Slide 20: UA Curve Fitting
Samples of UA traces extracted and fit for testing
Fitting log-normal curve y = a +b * log(ec*t2)
Many different shapes
Not confident in optimizer

## Slide 21: UA Curve Fitting - Conclusion
Challenging to find a consistent curve fitting method that will work
If it could work consistently, we could extract extra insights from UA shapes:
  Bell curve or log-normal
  Width/height factor

## Slide 22
Text
Text
Text
Text

## Slide 23
Text
Text
Text