## Slide 1: Keeping Labor Safe – Phase 2
Machine Learning Outcomes
Lucas Finco

## Slide 2: Analysis Process (to date)
Update trace plotting to include mark-ups
Collection of Ground Truth
  193 traces as of 1/6/2020
Evaluation methods of algorithm established
Tuning of algorithm parameters

## Slide 3: UA Contraction Parameter Tuning
Two Parameters to tune
  Minimum time duration to be called a contraction
  Minimum rise in UA for UC threshold
  Alternate UC threshold using % of Max-Min UA
Three graphs provided:
  False Positive Rate vs. False Negative Rate
  Mean Square Error and Mean Absolute Error (of number of contractions)
  Precision vs. Recall

## Slide 4: UA Contraction Parameter – Minimum Duration
Recommended Value: 14s
Notes: 44min 55sec to run 41 cases

## Slide 5: UA Contraction Parameter – Minimum UA Gap
Recommended Value: 5.5

## Slide 6: UA FRI Component
Three ways to calculate FRI component
More than 5 UC’s
U/R Ratio > 4
More than x crossings of min-max value

## Slide 7: FHR Parameter Tuning
Parameter to tune
  Threshold for acceleration/deceleration
Three graphs provided:
  False Positive Rate vs. False Negative Rate
  Mean Square Error and Mean Absolute Error (count of accels)
  Precision vs. Recall

## Slide 8: FHR Acceleration – FHR Acceleration Gap
Recommended Value: 10

## Slide 9: FHR Late Recovery – Seconds after Contraction
Recommended Value: 3

## Slide 10: FHR Late Recovery – AUC after Contraction
Recommended Value: -100

## Slide 11: Advanced Plotting
Improved visualization with deltas and histograms

## Slide 12: Results: Excessive UA
Number of contractions exceeds 4 in 10 minutes
Accuracy: 85%
Sensitivity, Recall: 91%
Specificity, TNR: 69%
Precision, PPV: 89%
NPV: 75%
Excellent result, room for improvement
 | Predicted |  | 
Actual |  | 0 | 1
 | 0 | 36 | 16
 | 1 | 12 | 129

## Slide 13: Results: FHR Accelerations
FHR Accelerations = 0
Accuracy: 65%
Sensitivity, Recall: 82%
Specificity, TNR: 50%
Precision, PPV: 59%
NPV: 76%
Accel Count MAE: 1.49
Needs Improvement
 | Predicted |  | 
Actual |  | 0 | 1
 | 0 | 41 | 41
 | 1 | 13 | 59

## Slide 14: Results: FHR Baseline
FHR Baseline < 110 or >= 160
Accuracy: 95%
Sensitivity, Recall: 95%
Specificity, TNR: 66%
Precision, PPV: 99%
NPV: 22%
Very strong result, note only one zero, no improvement needed
 | Predicted |  | 
Actual |  | 0 | 1
 | 0 | 2 | 1
 | 1 | 9 | 142

## Slide 15: Results: FHR Variability
FHR Variability < 5 or > 15
Accuracy: 81%
Sensitivity, Recall: 98%
Specificity: 21%
Precision, PPV: 81%
NPV: 78%
Good result, better at predicting zeros, which is important, but missed 5 zeros
 | Predicted |  | 
Actual |  | 0 | 1
 | 0 | 16 | 18
 | 1 | 8 | 112

## Slide 16: Results: FHR Late Recovery
FHR Late Recovery Present
Accuracy: 77%
Sensitivity, Recall: 88%
Specificity: 50%
Precision, PPV: 82%
NPV: 60%
True Negative Rate: 50%
Good result, confident in prediction, but weak with actual zeros, and 2 false negatives
 | Predicted |  | 
Actual |  | 0 | 1
 | 0 | 17 | 25
 | 1 | 30 | 82

## Slide 17: Case Study: CaseID 26789953
Notes: Note FHR baseline tachycardia properly identified

## Slide 18: Case Study: CaseID 27034378
Notes: Note FHR baseline properly identified

## Slide 19: ToDo’s:
Use AUC method for late recoveries
Analysis Confidence Metric

## Slide 20: FRI Correlations to Outcomes
Notes: As of 1/7/2020

## Slide 21: FRI Correlations to Outcomes

## Slide 22: Recommendations
Collect ground truth at the source of the data
  When contractions happen
  When accelerations happen
  When late recoveries happen

## Slide 23: Appendix